## Data Summary
|  | Puffer Parameters | [5 empty] | FRIENDS Measurements | [5 empty] | Vape Info.
| Vape | Number of Puffs | Puff Time (s) | Total Puff Time (s) | Interpuff Time (s) | Puff Volume (mL) | Puff Flow Rate (mL/s) | Puffs Counted | Puff Count RE (puff-weighted) | Puff Count RE ± 95% CI (ratio estimator) | Total Puff Time (s) | Puff Time RE (puff-time-weighted) | Puff Time RE ± 95% CI (ratio estimator) | Thermistor Used (Y/N) | Thermistor Needed (Y/N) | Characteristics |  |  | n (sessions) | Σ Delivered Puffs | Σ Delivered Puff Time (s)
1 | Adjust Myflavor | 440 | 3 | 1320 | 18 | 120 | 40 | 440 | 0% | 0% | 1437.6144 | 8.91% | 0.5524% | Y | N | Closed System; Rechargeable | Tested By Brett |  | 11 | 440 | 1320
2 | Artery CL6 | 433 | 3 | 1299 | 18 | 120 | 40 | 403 | 6.928% | 9.951% | 1311.6472 | 9.196% | 10.11% | Y | N | Closed System; Rechargeable | Not updated in the website |  | 7 | 433 | 1299
3 | Breeze Pro | 400 | 3 | 1200 | 18 | 120 | 40 | 397 | 0.75% | 2.082% | 1241.0303 | 3.419% | 1.373% | Y | Y | Closed System; Non-rechargeable |  | ` | 5 | 400 | 1200
4 | EBCreate BC 5000 | 450 | 3 | 1350 | 18 | 120 | 40 | 446 | 0.8889% | 1.304% | 1358.2995 | 1.294% | 0.8055% | Y | Y | Closed System; Rechargeable |  |  | 5 | 450 | 1350
5 | Elf Bar BC5000 | 448 | 3 | 1344 | 18 | 120 | 40 | 448 | 0% | 0% | 1380.2581 | 2.884% | 1.585% | Y | Y | Closed System; Rechargeable |  |  | 6 | 448 | 1344
6 | Funky Republic ti7000 | 560 | 3 | 1680 | 18 | 120 | 40 | 558 | 0.7143% | 1.311% | 1794.0507 | 6.789% | 2.411% | Y | Y | Closed System; Rechargeable |  |  | 7 | 560 | 1680
7 | Fzzybar X20000 | 530 | 3 | 1590 | 18 | 120 | 40 | 490 | 7.547% | 12.52% | 1582.4889 | 6.862% | 7.627% | Y | Y | Closed System; Rechargeable |  |  | 5 | 530 | 1590
8 | Geek Bar Pulse | 400 | 3 | 1200 | 18 | 120 | 40 | 403 | 0.75% | 1.388% | 1252.7102 | 4.393% | 0.7203% | Y | Y | Closed System; Rechargeable |  |  | 5 | 400 | 1200
9 | Hyde IQ5000 | 536 | 3 | 1608 | 18 | 120 | 40 | 505 | 6.157% | 5.701% | 1450.9639 | 11.49% | 7.043% | Y | Y | Closed System; Rechargeable |  |  | 8 | 536 | 1608
10 | Hyde N-bar | 240 | 3 | 720 | 18 | 120 | 40 | 236 | 1.667% | 7.171% | 762.3135 | 8.956% | 9.379% | Y | Y | Closed System; Rechargeable |  |  | 3 | 240 | 720
11 | Hyde Rebel | 275 | 3 | 825 | 18 | 120 | 40 | 271 | 2.182% | 2.86% | 855.3565 | 3.68% | 2.968% | Y | Y | Closed System; Rechargeable |  |  | 5 | 275 | 825
12 | Juul | 185 | 3 | 555 | 18 | 120 | 40 | 185 | 0% | 0% | 550.9164 | 0.7358% | 1.713% | N | N | Pod; Non-refillable |  |  | 3 | 185 | 555
13 | Kangvape Speedy Beast X | 360 | 3 | 1080 | 18 | 120 | 40 | 357 | 1.389% | 2.562% | 1062.3376 | 3.305% | 3.521% | Y | N | Closed System; Rechargeable |  |  | 9 | 360 | 1080
15 | Lost Angel Sphere 35K | 480 | 3 | 1440 | 18 | 120 | 40 | 480 | 0% | 0% | 1481.2303 | 2.863% | 0.1482% | Y | N | Closed System; Rechargable |  |  | 6 | 480 | 1440
16 | Lost Mary MT15000 | 455 | 3 | 1365 | 18 | 120 | 40 | 449 | 2.637% | 1.561% | 1395.8214 | 2.586% | 2.101% | Y | Y | Closed System; Rechargeable |  |  | 6 | 455 | 1365
17 | Mr Fog Aura | 1049 | 3 | 3147 | 18 | 120 | 40 | 1013 | 5.338% | 6.745% | 3210.8442 | 10.81% | 6.592% | Y | Y | Closed System; Rechargeable |  |  | 11 | 1049 | 3147
18 | MR Fog Nova | 1000 | 3 | 3000 | 18 | 120 | 40 | 1004 | 0.6% | 0.5395% | 3249.2353 | 8.308% | 0.8124% | Y | Y | Closed System; Rechargeable |  |  | 25 | 1000 | 3000
19 | Njoy Ace | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 0% | N/A (n=1) | 285.3583 | 18.9% | N/A (n=1) | N | N | Pod; Non-refillable |  |  | 1 | 80 | 240
21 | North 5000 Disposable | 240 | 3 | 720 | 18 | 120 | 40 | 240 | 0% | 0% | 772.4195 | 7.28% | 0.6853% | Y | N | Closed System; Rechargeable |  |  | 6 | 240 | 720
22 | Offstamp SW 9000 | 600 | 3 | 1800 | 18 | 120 | 40 | 604 | 0.6667% | 0.8218% | 1952.5325 | 8.474% | 0.9089% | Y | Y | Pod; Non-refillable |  |  | 15 | 600 | 1800
23 | Offstamp X cube | 330 | 3 | 990 | 18 | 120 | 40 | 293 | 11.21% | 8.993% | 880.5201 | 14.66% | 7.911% | Y | Y | Pod; Non-refillable |  |  | 6 | 330 | 990
24 | RAZ DC25000 | 502 | 3 | 1506 | 18 | 120 | 40 | 497 | 1.394% | 2.291% | 1555.7735 | 3.982% | 1.176% | Y | Y | Closed System; Rechargeable |  |  | 7 | 502 | 1506
25 | Tyson Heavyweight | 600 | 3 | 1800 | 18 | 120 | 40 | 601 | 1.167% | 0.7149% | 1951.3244 | 8.407% | 0.9329% | Y | Y | Closed System; Rechargeable |  |  | 15 | 600 | 1800
26 | Viho TRX 50k | 826 | 3 | 2478 | 18 | 120 | 40 | 818 | 0.9685% | 1.414% | 2636.1054 | 6.38% | 1.484% | Y | Y | Closed System; Rechargeable |  |  | 8 | 826 | 2478
27 | Fogger Switch Pro | 400 | 3 | 1200 | 18 | 120 | 40 | 260 | 35% | 58.51% | 797.3545 | 33.55% | 60.01% | Y | Y | Pod; Non-refillable |  |  | 4 | 400 | 1200
28 | Juul 2 | 240 | 3 | 720 | 18 | 120 | 40 | 1 | 99.58% | 1.793% | 0.4032 | 99.94% | 0.2409% | N | N | Pod; Non-refillable |  |  | 3 | 240 | 720
29 | KangerTech SUBOX | 650 | 3 | 1950 | 18 | 120 | 40 | 150 | 76.92% | 5.279% | 385.3958 | 80.24% | 11.53% | Y | N | Closed System; Rechargeable |  |  | 5 | 650 | 1950
30 | LostMary OS5000 | 560 | 3 | 1680 | 18 | 120 | 40 | 326 | 41.79% | 15.76% | 788.0675 | 53.09% | 22.34% | Y | N | Closed System; Rechargeable |  |  | 7 | 560 | 1680

## CI Method
Confidence intervals for FRIENDS accuracy: method comparison
Old = unweighted t-CI over session-level ¦RE¦ (previous Data Summary K/N). New = puff-weighted ratio estimator with sessions as clusters (current Data Summary K/N).
Vape | n sessions | Σ puffs delivered | Min session size | Max session size | OLD count ¦RE¦ mean | OLD ±95% CI | OLD CI lower bound | NEW count WMAPE | NEW ±95% CI | Signed count bias | Signed ±95% CI | OLD time ¦RE¦ mean | OLD ±95% CI | NEW time WMAPE | NEW ±95% CI | Signed time bias | Signed ±95% CI
Adjust Myflavor | 11 | 440 | 40 | 40 | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 8.91% | 0.5524% | 8.91% | 0.5524% | 8.91% | 0.5524%
Artery CL6 | 7 | 433 | 16 | 140 | 10.77% | 15.72% | -4.954% | 6.928% | 9.951% | -6.928% | 9.951% | 13.46% | 11.85% | 9.196% | 10.11% | 0.9736% | 11.26%
Breeze Pro | 5 | 400 | 80 | 80 | 0.75% | 2.082% | -1.332% | 0.75% | 2.082% | -0.75% | 2.082% | 3.419% | 1.373% | 3.419% | 1.373% | 3.419% | 1.373%
EBCreate BC 5000 | 5 | 450 | 70 | 100 | 1.021% | 1.457% | -0.4358% | 0.8889% | 1.304% | -0.8889% | 1.304% | 1.313% | 0.8851% | 1.294% | 0.8055% | 0.6148% | 1.666%
Elf Bar BC5000 | 6 | 448 | 48 | 80 | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 2.935% | 1.505% | 2.884% | 1.585% | 2.698% | 1.992%
Funky Republic ti7000 | 7 | 560 | 80 | 80 | 0.7143% | 1.311% | -0.5966% | 0.7143% | 1.311% | -0.3571% | 1.449% | 6.789% | 2.411% | 6.789% | 2.411% | 6.789% | 2.411%
Fzzybar X20000 | 5 | 530 | 30 | 140 | 6.433% | 11.01% | -4.581% | 7.547% | 12.52% | -7.547% | 12.52% | 6.645% | 6.735% | 6.862% | 7.627% | -0.4724% | 12.97%
Geek Bar Pulse | 5 | 400 | 80 | 80 | 0.75% | 1.388% | -0.6382% | 0.75% | 1.388% | 0.75% | 1.388% | 4.393% | 0.7203% | 4.393% | 0.7203% | 4.393% | 0.7203%
Hyde IQ5000 | 8 | 536 | 16 | 80 | 5.759% | 5.228% | 0.5306% | 6.157% | 5.701% | -5.784% | 6.059% | 12.57% | 7.15% | 11.49% | 7.043% | -9.766% | 9.076%
Hyde N-bar | 3 | 240 | 80 | 80 | 1.667% | 7.171% | -5.504% | 1.667% | 7.171% | -1.667% | 7.171% | 8.956% | 9.379% | 8.956% | 9.379% | 5.877% | 22.6%
Hyde Rebel | 5 | 275 | 35 | 80 | 2.893% | 3.017% | -0.1246% | 2.182% | 2.86% | -1.455% | 3.648% | 4.027% | 3.796% | 3.68% | 2.968% | 3.68% | 2.968%
Juul | 3 | 185 | 25 | 80 | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 0.7997% | 1.387% | 0.7358% | 1.713% | -0.7358% | 1.713%
Kangvape Speedy Beast X | 9 | 360 | 40 | 40 | 1.389% | 2.562% | -1.173% | 1.389% | 2.562% | -0.8333% | 2.718% | 3.305% | 3.521% | 3.305% | 3.521% | -1.635% | 4.228%
Lost Angel Sphere 35K | 6 | 480 | 80 | 80 | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 2.863% | 0.1482% | 2.863% | 0.1482% | 2.863% | 0.1482%
Lost Mary MT15000 | 6 | 455 | 55 | 80 | 2.689% | 1.517% | 1.172% | 2.637% | 1.561% | -1.319% | 2.91% | 2.797% | 2.411% | 2.586% | 2.101% | 2.258% | 2.546%
Mr Fog Aura | 11 | 1049 | 39 | 200 | 11.16% | 12.99% | -1.833% | 5.338% | 6.745% | -3.432% | 7.132% | 15.44% | 13.7% | 10.81% | 6.592% | 2.029% | 8.764%
MR Fog Nova | 25 | 1000 | 40 | 40 | 0.6% | 0.5395% | 0.06048% | 0.6% | 0.5395% | 0.4% | 0.5715% | 8.308% | 0.8124% | 8.308% | 0.8124% | 8.308% | 0.8124%
Njoy Ace | 1 | 80 | 80 | 80 | 0% | N/A | N/A | 0% | N/A (n=1) | 0% | N/A | 18.9% | N/A | 18.9% | N/A (n=1) | 18.9% | N/A
North 5000 Disposable | 6 | 240 | 40 | 40 | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 7.28% | 0.6853% | 7.28% | 0.6853% | 7.28% | 0.6853%
Offstamp SW 9000 | 15 | 600 | 40 | 40 | 0.6667% | 0.8218% | -0.1552% | 0.6667% | 0.8218% | 0.6667% | 0.8218% | 8.474% | 0.9089% | 8.474% | 0.9089% | 8.474% | 0.9089%
Offstamp X cube | 6 | 330 | 40 | 70 | 10% | 9.832% | 0.1679% | 11.21% | 8.993% | -11.21% | 8.993% | 13.97% | 7.96% | 14.66% | 7.911% | -11.06% | 12.7%
RAZ DC25000 | 7 | 502 | 22 | 80 | 1.25% | 2.111% | -0.8607% | 1.394% | 2.291% | -0.996% | 2.519% | 3.885% | 1.104% | 3.982% | 1.176% | 3.305% | 2.628%
Tyson Heavyweight | 15 | 600 | 40 | 40 | 1.167% | 0.7149% | 0.4517% | 1.167% | 0.7149% | 0.1667% | 0.9743% | 8.407% | 0.9329% | 8.407% | 0.9329% | 8.407% | 0.9329%
Viho TRX 50k | 8 | 826 | 40 | 176 | 0.9146% | 1.607% | -0.6921% | 0.9685% | 1.414% | -0.9685% | 1.414% | 6.413% | 1.748% | 6.38% | 1.484% | 6.38% | 1.484%
Fogger Switch Pro | 4 | 400 | 100 | 100 | 35% | 58.51% | -23.51% | 35% | 58.51% | -35% | 58.51% | 33.55% | 60.01% | 33.55% | 60.01% | -33.55% | 60.01%
Juul 2 | 3 | 240 | 80 | 80 | 99.58% | 1.793% | 97.79% | 99.58% | 1.793% | -99.58% | 1.793% | 99.94% | 0.2409% | 99.94% | 0.2409% | -99.94% | 0.2409%
KangerTech SUBOX | 5 | 650 | 120 | 140 | 76.77% | 5.083% | 71.68% | 76.92% | 5.279% | -76.92% | 5.279% | 79.92% | 10.95% | 80.24% | 11.53% | -80.24% | 11.53%
LostMary OS5000 | 7 | 560 | 80 | 80 | 41.79% | 15.76% | 26.02% | 41.79% | 15.76% | -41.79% | 15.76% | 53.09% | 22.34% | 53.09% | 22.34% | -53.09% | 22.34%
Why the old formula was wrong
1. Unit-of-analysis mismatch. The old CI treats every puff session as one equally weighted observation, but sessions within a vape range from 16 to 200 delivered puffs. A session error rate estimated from 16 puffs is far noisier than one from 200 puffs, so the session-level ¦RE¦ values are heteroscedastic and the unweighted mean/STDEV is neither efficient nor correctly scaled.
2. Different total puffs per vape. Because the estimand is a per-puff error rate, the natural denominator is total puffs delivered, not number of sessions. The number of sessions is an administrative artifact of how testing was batched.
3. Symmetric t-CI on a non-negative folded quantity. ¦RE¦ >= 0, so a symmetric interval can (and for 12 of 28 vapes did) extend below zero. See column H.
New method: Cochran ratio estimator, sessions as clusters
For vape v with sessions j = 1..n: x_j = delivered puffs (or delivered puff time), u_j = ¦counted - delivered¦.
Point estimate R = Σ u_j / Σ x_j (puff-weighted mean absolute relative error, WMAPE)
Variance v(R) = n * Σ (u_j - R*x_j)^2 / [ (n-1) * (Σ x_j)^2 ]
95% CI R ± t(0.975, n-1) * SQRT(v(R))
This weights each session by its own size, makes no distributional assumption about within-session errors, and is robust to correlation of errors within a session. It is the standard survey-sampling variance for a ratio of totals (Cochran 1977, Sampling Techniques, 3rd ed., Sec. 6.3).
Assumptions and limitations
a) Sessions within a vape are treated as independent replicate clusters. If sessions were run back-to-back on one coil, this is optimistic (coil ageing induces between-session drift that is not exchangeable).
b) n = 1 (Njoy Ace) admits no session-level variance estimate; K/N return "N/A (n=1)". A puff-level Wilson score interval on the detection rate (80/80 detected -> 95.4%-100%) is the only alternative, but see (c).
c) A puff-level binomial/Wilson interval is NOT recommended as the primary method. Comparing the cluster variance to the binomial variance gives design effects of roughly 3x to 60x on this data (e.g. Mr Fog Aura ~32x, Fogger Switch Pro ~59x). Puff detection failures arrive in bursts, so a binomial CI would be several-fold too narrow.
d) A CI of exactly 0.00% means every session showed identical error (usually zero error at every puff). Report these as "no error observed" with a rule-of-three upper bound 3/N (N = total puffs), inflated by the design effect, rather than as a zero-width interval.
e) ¦RE¦ is not scale-invariant across sessions of different length. If detection errors were independent across puffs, E¦RE¦ would fall as 1/sqrt(session length); if errors are systematic, E¦RE¦ is constant. In this dataset errors are dominated by systematic undercounting, so weighting by x_j is appropriate - but this should be stated, and the signed bias columns (K, Q) reported alongside as the Bland-Altman bias term.
f) Across-vape pooling is a separate question. Do not average the 28 vape-level values as if they were 28 independent observations of one quantity; vapes differ systematically (see Juul 2, KangerTech SUBOX). Use a random-effects model, or report the distribution (median and range) of vape-level WMAPE.
Other corrections applied to Data Summary
- M19 (MR Fog Nova, puff time RE) referenced 'All Data'!M126:M136, which are Mr Fog Aura rows. Corrected to M137:M161 to match N19 and J19/K19.
- M23, N23, M31, N31 referenced an external workbook ('[1]All Data'). Re-pointed to the local 'All Data' sheet; cached values matched, so results are unchanged.
- Columns U:W added to Data Summary (n sessions, Σ delivered puffs, Σ delivered puff time) as explicit, auditable denominators.
- Columns J and M now hold the puff-weighted point estimate, because a CI must belong to the same estimator as the point estimate it accompanies. Old values are preserved in columns F and M of this sheet.

## All Data
Vape | Puff Session | Number of Puffs | Actual Puff Time (s) | Total Puff Time (s) | Interpuff Time (s) | Puff Volume (mL) | Puff Flow Rate (mL/s) | Puffs Counted | Recorded Puff Time (s) | Puff Count RE (per session) | Puff Time RE (per session)
Elf Bar BC5000 | 1 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 238.7464 | 0 | 0.0052233
| 2 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 244.4862 | 0 | 0.018693
| 3 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 250.0572 | 0 | 0.041905
| 4 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 249.418 | 0 | 0.039242
| 5 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 248.3032 | 0 | 0.034597
| 6 | 48 | 3 | 144 | 18 | 120 | 40 | 48 | 149.2471 | 0 | 0.036438
| Overall | 448 | - | 1344 | - | - | - | 448 | 1380.2581
Geek Bar Pulse | 1 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 249.9346 | 0 | 0.041394
| 2 | 80 | 3 | 240 | 18 | 120 | 40 | 82 | 252.4966 | 0.025 | 0.052069
| 3 | 80 | 3 | 240 | 18 | 120 | 40 | 81 | 251.0437 | 0.0125 | 0.046015
| 4 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 250.5128 | 0 | 0.043803
| 5 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 248.7225 | 0 | 0.036344
| Overall | 400 | - | 1200 | - | - | - | 403 | 1252.7102
RAZ DC25000 | 1 | 80 | 3 | 240 | 18 | 120 | 40 | 81 | 253.9351 | 0.0125 | 0.058063
| 2 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 250.8108 | 0 | 0.045045
| 3 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 250.3718 | 0 | 0.043216
| 4 | 80 | 3 | 240 | 18 | 120 | 40 | 79 | 247.8016 | 0.0125 | 0.032507
| 5 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 249.9447 | 0 | 0.041436
| 6 | 80 | 3 | 240 | 18 | 120 | 40 | 75 | 234.9017 | 0.0625 | 0.021243
| 7 | 22 | 3 | 66 | 18 | 120 | 40 | 22 | 68.0078 | 0 | 0.030421
| Overall | 502 | - | 1506 | - | - | - | 497 | 1555.7735
Lost Mary MT15000 | 1 | 80 | 3 | 240 | 18 | 120 | 40 | 79 | 242.5057 | 0.0125 | 0.01044
| 2 | 80 | 3 | 240 | 18 | 120 | 40 | 78 | 244.0737 | 0.025 | 0.016974
| 3 | 80 | 3 | 240 | 18 | 120 | 40 | 81 | 250.9333 | 0.0125 | 0.045555
| 4 | 80 | 3 | 240 | 18 | 120 | 40 | 76 | 237.7629 | 0.05 | 0.0093212
| 5 | 80 | 3 | 240 | 18 | 120 | 40 | 78 | 244.5904 | 0.025 | 0.019127
| 6 | 55 | 3 | 165 | 18 | 120 | 40 | 57 | 175.9554 | 0.036364 | 0.066396
| Overall | 455 | - | 1365 | - | - | - | 449 | 1395.8214
Funky Republic ti7000 | 1 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 250.7872 | 0 | 0.044947
| 2 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 254.519 | 0 | 0.060496
| 3 | 80 | 3 | 240 | 18 | 120 | 40 | 77 | 245.8975 | 0.0375 | 0.024573
| 4 | 80 | 3 | 240 | 18 | 120 | 40 | 81 | 261.9976 | 0.0125 | 0.091657
| 5 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 261.9183 | 0 | 0.091326
| 6 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 261.5835 | 0 | 0.089931
| 7 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 257.3476 | 0 | 0.072282
| Overall | 560 | - | 1680 | - | - | - | 558 | 1794.0507
Breeze Pro | 1 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 249.3192 | 0 | 0.03883
| 2 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 248.4511 | 0 | 0.035213
| 3 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 249.44 | 0 | 0.039333
| 4 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 250.2273 | 0 | 0.042614
| 5 | 80 | 3 | 240 | 18 | 120 | 40 | 77 | 243.5927 | 0.0375 | 0.01497
| Overall | 400 | - | 1200 | - | - | - | 397 | 1241.0303
Nord 2 | 1 | 30 | 3 | 87.06 | 18 | - | - | 30 | 85.0256
| Overall | 30 | - | 87.06 | - | - | - | 30 | 85.0256
Lost Angel Sphere 35K | 1 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 247.2653 | 0 | 0.030272
| 2 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 246.96 | 0 | 0.029
| 3 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 246.7951 | 0 | 0.028313
| 4 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 247.0036 | 0 | 0.029182
| 5 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 246.2537 | 0 | 0.026057
| 6 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 246.9526 | 0 | 0.028969
| Overall | 480 | - | 1440 | - | - | - | 480 | 1481.2303
Juul | 1 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 237.0756 | 0 | 0.012185
| 2 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 239.6019 | 0 | 0.0016588
| 3 | 25 | 3 | 75 | 18 | 120 | 40 | 25 | 74.2389 | 0 | 0.010148
| Overall | 185 | - | 555 | - | - | - | 185 | 550.9164
Njoy Ace | 1 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 285.3583
| Overall | 80 | - | 240 | - | - | - | 80 | 285.3583
Logic Pro | 1 | 30 | 3 | 94.6828 | 18 | - | - | 30 | 87.3248
| Overall | 30 | - | 94.6828 | - | - | - | 30 | 87.3248
Adjust Myflavor | 1 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 130.3667 | 0 | 0.086389
| 2 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 130.3418 | 0 | 0.086182
| 3 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 130.648 | 0 | 0.088733
| 4 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 131.5078 | 0 | 0.095898
| 5 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 128.0169 | 0 | 0.066807
| 6 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 130.9131 | 0 | 0.090942
| 7 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 130.8764 | 0 | 0.090637
| 8 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 130.8097 | 0 | 0.090081
| 9 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 131.7166 | 0 | 0.097638
| 10 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 131.2452 | 0 | 0.09371
| 11 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 131.1722 | 0 | 0.093102
| Overall | 440 | - | 1320 | - | - | - | 440 | 1437.6144
Artery CL6 | 1 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 129.4749 | 0 | 0.078957
| 2 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 260.7989 | 0 | 0.086662
| 3 | 16 | 3 | 48 | 18 | 120 | 40 | 12 | 38.5496 | 0.25 | 0.19688
| 4 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 263.7534 | 0 | 0.098973
| 5 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 129.3647 | 0 | 0.078039
| 6 | 37 | 3 | 111 | 18 | 120 | 40 | 21 | 67.044 | 0.43243 | 0.396
| 7 | 140 | 3 | 420 | 18 | 120 | 40 | 130 | 422.6617 | 0.071429 | 0.0063374
| Overall | 433 | - | 1299 | - | - | - | 403 | 1311.6472
EBCreate BC 5000 | 1 | 100 | 3 | 300 | 18 | 120 | 40 | 100 | 304.5612 | 0 | 0.015204
| 2 | 100 | 3 | 300 | 18 | 120 | 40 | 100 | 305.0452 | 0 | 0.016817
| 3 | 100 | 3 | 300 | 18 | 120 | 40 | 99 | 302.2399 | 0.01 | 0.0074663
| 4 | 80 | 3 | 240 | 18 | 120 | 40 | 79 | 241.0387 | 0.0125 | 0.0043279
| 5 | 70 | 3 | 210 | 18 | 120 | 40 | 68 | 205.4145 | 0.028571 | 0.021836
| Overall | 450 | - | 1350 | - | - | - | 446 | 1358.2995
Fzzybar X20000 | 1 | 130 | 3 | 390 | 18 | 120 | 40 | 129 | 416.1451 | 0.0076923 | 0.067039
| 2 | 100 | 3 | 300 | 18 | 120 | 40 | 93 | 300.3535 | 0.07 | 0.0011783
| 3 | 30 | 3 | 90 | 18 | 120 | 40 | 30 | 96.4904 | 0 | 0.072116
| 4 | 140 | 3 | 420 | 18 | 120 | 40 | 136 | 437.8052 | 0.028571 | 0.042393
| 5 | 130 | 3 | 390 | 18 | 120 | 40 | 102 | 331.6947 | 0.21538 | 0.1495
| Overall | 530 | - | 1590 | - | - | - | 490 | 1582.4889
Hyde IQ5000 | 1 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 248.9808 | 0 | 0.03742
| 2 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 217.4221 | 0 | 0.094075
| 3 | 80 | 3 | 240 | 18 | 120 | 40 | 67 | 192.0866 | 0.1625 | 0.19964
| 4 | 80 | 3 | 240 | 18 | 120 | 40 | 72 | 204.1822 | 0.1 | 0.14924
| 6 | 16 | 3 | 48 | 18 | 120 | 40 | 16 | 40.5957 | 0 | 0.15426
| 7 | 70 | 3 | 210 | 18 | 120 | 40 | 64 | 193.9469 | 0.085714 | 0.076443
| 8 | 80 | 3 | 240 | 18 | 120 | 40 | 81 | 244.8702 | 0.0125 | 0.020293
| 9 | 50 | 3 | 150 | 18 | 120 | 40 | 45 | 108.8794 | 0.1 | 0.27414
| Overall | 536 | - | 1608 | - | - | - | 505 | 1450.9639
Hyde N-bar | 1 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 267.6167 | 0 | 0.11507
| 2 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 265.7834 | 0 | 0.10743
| 3 | 80 | 3 | 240 | 18 | 120 | 40 | 76 | 228.9134 | 0.05 | 0.046194
| Overall | 240 | - | 720 | - | - | - | 236 | 762.3135
Hyde REBEL | 1 | 80 | 3 | 240 | 18 | 120 | 40 | 81 | 249.6478 | 0.0125 | 0.040199
| 2 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 244.9321 | 0 | 0.02055
| 3 | 40 | 3 | 120 | 18 | 120 | 40 | 39 | 120.0784 | 0.025 | 0.00065333
| 4 | 40 | 3 | 120 | 18 | 120 | 40 | 38 | 128.0449 | 0.05 | 0.067041
| 5 | 35 | 3 | 105 | 18 | 120 | 40 | 33 | 112.6533 | 0.057143 | 0.072889
| Overall | 275 | - | 825 | - | - | - | 271 | 855.3565
Kangvape Speedy Beast X | 1 | 40 | 3 | 120 | 18 | 120 | 40 | 36 | 108.9191 | 0.1 | 0.092341
| 2 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 104.4049 | 0 | 0.12996
| 3 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 120.853 | 0 | 0.0071083
| 4 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 120.8495 | 0 | 0.0070792
| 5 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 120.6962 | 0 | 0.0058017
| 6 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 121.2251 | 0 | 0.010209
| 7 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 120.5273 | 0 | 0.0043942
| 8 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 121.519 | 0 | 0.012658
| 9 | 40 | 3 | 120 | 18 | 120 | 40 | 41 | 123.3435 | 0.025 | 0.027863
| Overall | 360 | - | 1080 | - | - | - | 357 | 1062.3376
Mr Fog Aura | 1 | 39 | 3 | 117 | 18 | 120 | 40 | 30 | 96.4425 | 0.23077 | 0.17571
| 2 | 40 | 3 | 120 | 18 | 120 | 40 | 14 | 29.4417 | 0.65 | 0.75465
| 3 | 40 | 3 | 120 | 18 | 120 | 40 | 41 | 124.4651 | 0.025 | 0.037209
| 4 | 40 | 3 | 120 | 18 | 120 | 40 | 36 | 113.2754 | 0.1 | 0.056038
| 5 | 40 | 3 | 120 | 18 | 120 | 40 | 41 | 132.7363 | 0.025 | 0.10614
| 6 | 160 | 3 | 480 | 18 | 120 | 40 | 159 | 512.8645 | 0.00625 | 0.068468
| 7 | 40 | 3 | 120 | 18 | 120 | 40 | 34 | 99.7351 | 0.15 | 0.16887
| 8 | 200 | 3 | 600 | 18 | 120 | 40 | 200 | 638.1393 | 0 | 0.063566
| 9 | 200 | 3 | 600 | 18 | 120 | 40 | 207 | 654.0942 | 0.035 | 0.090157
| 10 | 190 | 3 | 570 | 18 | 120 | 40 | 191 | 610.5276 | 0.0052632 | 0.071101
| 11 | 60 | 3 | 180 | 18 | 120 | 40 | 60 | 199.1225 | 0 | 0.10624
| Overall | 1049 | - | 3147 | - | - | - | 1013 | 3210.8442
MR Fog Nova | 1 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 128.8517 | 0 | 0.073764
| 2 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 131.687 | 0 | 0.097392
| 3 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 129.8516 | 0 | 0.082097
| 4 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 130.7159 | 0 | 0.089299
| 5 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 132.1853 | 0 | 0.10154
| 6 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 131.5598 | 0 | 0.096332
| 7 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 129.3144 | 0 | 0.07762
| 8 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 122.0562 | 0 | 0.017135
| 9 | 40 | 3 | 120 | 18 | 120 | 40 | 41 | 131.4685 | 0.025 | 0.095571
| 10 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 129.9673 | 0 | 0.083061
| 11 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 130.2103 | 0 | 0.085086
| 12 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 129.744 | 0 | 0.0812
| 13 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 131.1165 | 0 | 0.092637
| 14 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 124.9706 | 0 | 0.041422
| 15 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 131.394 | 0 | 0.09495
| 16 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 131.0166 | 0 | 0.091805
| 17 | 40 | 3 | 120 | 18 | 120 | 40 | 41 | 132.8421 | 0.025 | 0.10702
| 18 | 40 | 3 | 120 | 18 | 120 | 40 | 42 | 133.2645 | 0.05 | 0.11054
| 19 | 40 | 3 | 120 | 18 | 120 | 40 | 39 | 127.5121 | 0.025 | 0.062601
| 20 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 129.6283 | 0 | 0.080236
| 21 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 129.2453 | 0 | 0.077044
| 22 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 129.9063 | 0 | 0.082552
| 23 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 129.8921 | 0 | 0.082434
| 24 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 129.8616 | 0 | 0.08218
| 25 | 40 | 3 | 120 | 18 | 120 | 40 | 41 | 130.9733 | 0.025 | 0.091444
| Overall | 1000 | - | 3000 | - | - | - | 1004 | 3249.2353
North 5000 Disposable | 1 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 128.6867 | 0 | 0.072389
| 2 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 129.147 | 0 | 0.076225
| 3 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 129.8414 | 0 | 0.082012
| 4 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 129.0301 | 0 | 0.075251
| 5 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 127.7064 | 0 | 0.06422
| 6 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 128.0079 | 0 | 0.066733
| Overall | 240 | - | 720 | - | - | - | 240 | 772.4195
Offstamp SW 9000 | 1 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 126.7972 | 0 | 0.056643
| 2 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 127.1815 | 0 | 0.059846
| 3 | 40 | 3 | 120 | 18 | 120 | 40 | 42 | 134.522 | 0.05 | 0.12102
| 4 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 130.1859 | 0 | 0.084882
| 5 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 131.1643 | 0 | 0.093036
| 6 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 132.0304 | 0 | 0.10025
| 7 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 130.5487 | 0 | 0.087906
| 8 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 128.2062 | 0 | 0.068385
| 9 | 40 | 3 | 120 | 18 | 120 | 40 | 41 | 129.4842 | 0.025 | 0.079035
| 10 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 129.2707 | 0 | 0.077256
| 11 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 131.4285 | 0 | 0.095238
| 12 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 128.903 | 0 | 0.074192
| 13 | 40 | 3 | 120 | 18 | 120 | 40 | 41 | 131.1328 | 0.025 | 0.092773
| 14 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 130.4423 | 0 | 0.087019
| 15 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 131.2348 | 0 | 0.093623
| Overall | 600 | - | 1800 | - | - | - | 604 | 1952.5325
Tyson Heavyweight | 1 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 129.4886 | 0 | 0.079072
| 2 | 40 | 3 | 120 | 18 | 120 | 40 | 41 | 129.5724 | 0.025 | 0.07977
| 3 | 40 | 3 | 120 | 18 | 120 | 40 | 39 | 126.7116 | 0.025 | 0.05593
| 4 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 130.0612 | 0 | 0.083843
| 5 | 40 | 3 | 120 | 18 | 120 | 40 | 39 | 126.8539 | 0.025 | 0.057116
| 6 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 129.7779 | 0 | 0.081482
| 7 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 127.7146 | 0 | 0.064288
| 8 | 40 | 3 | 120 | 18 | 120 | 40 | 41 | 132.9579 | 0.025 | 0.10798
| 9 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 130.7734 | 0 | 0.089778
| 10 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 132.3587 | 0 | 0.10299
| 11 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 130.459 | 0 | 0.087158
| 12 | 40 | 3 | 120 | 18 | 120 | 40 | 41 | 131.6567 | 0.025 | 0.097139
| 13 | 40 | 3 | 120 | 18 | 120 | 40 | 39 | 128.3914 | 0.025 | 0.069928
| 14 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 132.0504 | 0 | 0.10042
| 15 | 40 | 3 | 120 | 18 | 120 | 40 | 41 | 132.4967 | 0.025 | 0.10414
| Overall | 600 | - | 1800 | - | - | - | 601 | 1951.3244
Viho TRX 50k | 1 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 129.6214 | 0 | 0.080178
| 2 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 259.5422 | 0 | 0.081426
| 3 | 80 | 3 | 240 | 18 | 120 | 40 | 80 | 258.3287 | 0 | 0.07637
| 4 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 126.9435 | 0 | 0.057862
| 5 | 160 | 3 | 480 | 18 | 120 | 40 | 160 | 512.2895 | 0 | 0.06727
| 6 | 160 | 3 | 480 | 18 | 120 | 40 | 159 | 514.5016 | 0.00625 | 0.071878
| 7 | 90 | 3 | 270 | 18 | 120 | 40 | 85 | 274.5489 | 0.055556 | 0.016848
| 8 | 176 | 3 | 528 | 18 | 120 | 40 | 174 | 560.3296 | 0.011364 | 0.06123
| Overall | 826 | - | 2478 | - | - | - | 818 | 2636.1054
Fogger Switch Pro | 1 | 100 | 3 | 300 | 18 | 120 | 40 | 97 | 299.0495 | 0.03 | 0.0031683
| 2 | 100 | 3 | 300 | 18 | 120 | 40 | 67 | 204.9101 | 0.33 | 0.31697
| 3 | 100 | 3 | 300 | 18 | 120 | 40 | 83 | 253.6615 | 0.17 | 0.15446
| 4 | 100 | 3 | 300 | 18 | 120 | 40 | 13 | 39.7334 | 0.87 | 0.86756
| Overall | 400 | - | 1200 | - | - | - | 260 | 797.3545
Juul 2 | 1 | 80 | 3 | 240 | 18 | 120 | 40 | 1 | 0.4032 | 0.9875 | 0.99832
| 2 | 80 | 3 | 240 | 18 | 120 | 40 | 0 | 0 | 1 | 1
| 3 | 80 | 3 | 240 | 18 | 120 | 40 | 0 | 0 | 1 | 1
| Overall | 240 | - | 720 | - | - | - | 1
KangerTech SUBOX | 1 | 140 | 3 | 420 | 18 | 120 | 40 | 23 | 18.7398 | 0.83571 | 0.95538
| 2 | 120 | 3 | 360 | 18 | 120 | 40 | 32 | 90.494 | 0.73333 | 0.74863
| 3 | 130 | 3 | 390 | 18 | 120 | 40 | 30 | 88.5408 | 0.76923 | 0.77297
| 4 | 130 | 3 | 390 | 18 | 120 | 40 | 34 | 98.6325 | 0.73846 | 0.7471
| 5 | 130 | 3 | 390 | 18 | 120 | 40 | 31 | 88.9887 | 0.76154 | 0.77182
| Overall | 650 | - | 1950 | - | - | - | 150
LostMary OS5000 | 1 | 80 | 3 | 240 | 18 | 120 | 40 | 21 | 17.2341 | 0.7375 | 0.92819
| 2 | 80 | 3 | 240 | 18 | 120 | 40 | 53 | 134.1445 | 0.3375 | 0.44106
| 3 | 80 | 3 | 240 | 18 | 120 | 40 | 53 | 142.3407 | 0.3375 | 0.40691
| 4 | 80 | 3 | 240 | 18 | 120 | 40 | 55 | 148.405 | 0.3125 | 0.38165
| 5 | 80 | 3 | 240 | 18 | 120 | 40 | 34 | 40.4091 | 0.575 | 0.83163
| 6 | 80 | 3 | 240 | 18 | 120 | 40 | 57 | 151.642 | 0.2875 | 0.36816
| 7 | 80 | 3 | 240 | 18 | 120 | 40 | 53 | 153.8921 | 0.3375 | 0.35878
| Overall | 560 | - | 1680 | - | - | - | 326 | 788.0675
Offstamp X cube | 1 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 128.7469 | 0 | 0.072891
| 2 | 40 | 3 | 120 | 18 | 120 | 40 | 40 | 129.0558 | 0 | 0.075465
| 3 | 70 | 3 | 210 | 18 | 120 | 40 | 63 | 188.1083 | 0.1 | 0.10425
| 4 | 60 | 3 | 180 | 18 | 120 | 40 | 55 | 158.4368 | 0.083333 | 0.1198
| 5 | 60 | 3 | 180 | 18 | 120 | 40 | 48 | 142.4346 | 0.2 | 0.2087
| 6 | 60 | 3 | 180 | 18 | 120 | 40 | 47 | 133.7377 | 0.21667 | 0.25701
| Overall | 330 | - | 990 | - | - | - | 293 | 880.5201
Note on C1: Brett Fassler: Each puff session entails 80 puffs followed by a recharge. The final test may have less than 80 puffs if the vape died completely during the test.
Note on C2: Brett Fassler: Each puff session entails 80 puffs followed by a recharge. The final test may have less than 80 puffs if the vape died completely during the test.