## Per_Participant_Durations
Participant | Camera Avg Puff Duration | FRIENDS Avg Puff Duration | Avg Hit Duration (Camera) | Avg Hit Duration (FRIENDS)
FRIENDS2076 | 2.776 | 1.893 | 2.776 | 1.913
FRIENDS2080 | 1.156 | 1.209 | 1.161 | 1.209
FRIENDS2081 | 2.447 | 2.07 | 2.447 | 2.07
FRIENDS2091 | 3.211 | 2.725 | 3.211 | 2.744
FRIENDS2094 | 3.171 | 1.887 | 3.64 | 2.691
FRIENDS2112 | 2.924 | 2.462 | 2.924 | 2.462
FRIENDS2136 | 2.005 | 1.805 | 1.995 | 1.805
FRIENDS2137 | 1.702 | 2.16 | 1.684 | 1.66
FRIENDS2138 | 2.262 | 1.792 | 2.262 | 1.792
FRIENDS2150 | 3.024 | 2.753 | 3.024 | 2.753
FRIENDS2160 | 2.321 | 1.621 | 2.32 | 1.483
FRIENDS2164 | 3.309 | 2.927 | 3.309 | 2.927
FRIENDS2173 | 2.282 | 1.524 | 2.282 | 1.436
FRIENDS2191 | 1.636 | 1.286 | 1.636 | 1.299
FRIENDS2192 | 1.77 | 1.819 | 1.77 | 1.666
FRIENDS2202 | 3.653 | 2.944 | 3.653 | 2.944
FRIENDS2203 | 2.49 | 1.791 | 2.49 | 1.791
FRIENDS2209 | 1.589 | 1.088 | 1.589 | 1.111
FRIENDS2219 | 2.478 | 1.938 | 2.516 | 1.938
FRIENDS2221 | 2.444 | 1.925 | 2.444 | 2.106
FRIENDS2226 | 2.655 | 2.298 | 2.66 | 2.419
FRIENDS2236 | 4.833 | 2.793 | 4.833 | 2.793

## Summary
Duration definition | n | Camera mean (s) | Camera SD (s) | FRIENDS mean (s) | FRIENDS SD (s) | Pearson r | Pearson p | BA mean bias (s) | BA bias 95% CI (+/-) | BA LoA lower (s) | BA LoA upper (s) | Proportional bias r | Proportional bias p
Hits only | 22 | 2.5739 | 0.8348 | 2.046 | 0.5852 | 0.869 | 1.5468e-07 | 0.5279 | 0.1934 | -0.327 | 1.3829 | 0.5912 | 0.0038

## Method
Note
Duration = Avg Hit Duration (Camera) vs Avg Hit Duration (FRIENDS): only puffs that were correctly matched between video and device (TP puffs). This is the only valid definition for a Bland-Altman agreement analysis, since each pair must be two measurements of the same event -- an "all puffs" comparison would mix in FRIENDS false-positive durations (no video counterpart) and camera missed-puff durations (no FRIENDS counterpart), conflating detection disagreement (already covered by the Script1 confusion matrix) with duration-measurement disagreement.
Computed fresh from Script1_Puff_Analysis_Reconciled_Confusion_Matrix.py (imported unmodified), not from any pre-existing summary spreadsheet.
Bland-Altman: bias = mean(Camera-FRIENDS); LoA = bias +/- 1.96*SD(diff); proportional bias = Pearson r between per-participant mean and difference.